## Summary
PCCT Scanner Qualification — Summary Dashboard
Fill in the tracker sheets. This summary auto-updates from status columns.
Gate | Total Criteria | Pass | Fail | In Progress | Gate Result
Gate 1 — Technical Prerequisites | 4
Gate 2 — Workflow Integration | 6
Gate 3 — Reproducibility | 4
Gate 4 — Bias & Agreement | 3
Advisory — Operational Checks | 6

## Gate 1 - Technical
Gate 1 — Technical & Image Quality Prerequisites
Required — Scanner not eligible without passing all items
ID | Criterion | Threshold | Method | Owner | Status | Evidence Link / Location | Evidence Description | Date Completed | Notes
1.1 | DICOM Compliance | Full DICOM 3.0; mandatory tags present | Automated tag validation on ingestion |  | Not started
1.2 | Contrast Timing | Peak aortic HU ≥ 300 in ≥ 90% of cases | Aortic ROI measurement across paired dataset |  | Not started
1.3 | Image Noise (SNR/CNR) | Non-inferior to reference CTA (±15%) | Uniform phantom or aortic ROI SD comparison |  | Not started
1.4 | Reconstruction Kernel | Soft-tissue kernel available; sharp kernel documented | Protocol review + DICOM header | Murali | Pass | https://elucidbio.sharepoint.com/:w:/r/Technology/Documents/IP/PCCT/PT142_combined_analysis_COMPLETE_FINAL_v1.0.docx?d=w0fe5c7e3207c4790930cac6308e7bde7&csf=1&web=1&e=mQv35o |  |  | Kernel family influences local vessel segment and patient level plaque characterization. Recon settings shoud be standardized for subsequent evaluations especially reproducibility and bias-agreement analyses

## Gate 2 - Workflow
Gate 2 — Workflow Integration Checks
Workflow step — Engineering remediation required before performance testing
ID | Criterion | Threshold | Method | Owner | Status | Evidence Link / Location | Evidence Description | Date Completed | Notes
2.1 | DICOM Ingestion | 100% successful ingestion; zero WIID assignment failures | Run full paired dataset through ingest pipeline |  | Not started
2.2 | Centerline / Vessel Tree | Extraction success rate ≥ 95%; major vessel coverage equivalent to reference | Compare vessel segment count and length distribution |  | Not started
2.3 | Lumen & Wall Initialization | Auto-initialization without manual override in ≥ 85% of segments | Log override frequency per scanner type |  | Not started
2.4 | Lumen & Wall Editing | Edit rate not significantly higher than reference CTA (95% CI overlap) | Track editing events per case; include cases Carolyn already completed |  | Not started
2.5 | Plaque Quantification | See Gate 3 — core quantitative criteria | Paired CTA vs. PCCT comparison |  | Not started
2.6 | Report Generation | All required fields populated; no missing-data warnings | ECR review in physician viewer and PDF reports for N ≥ 10 cases; use ones Carolyn already did |  | Not started

## Gate 3 - Reproducibility
Gate 3 — Core Quantitative Reproducibility (Paired CTA vs. PCCT)
Performance — wCV from un-edited output, manual centerline for N ≥ 30. All thresholds must pass.
ID | Criterion | Threshold | Method | Owner | Status | Evidence Link / Location | Evidence Description | Date Completed | Notes
3.1 | Lumen Volume & Vessel Length | ≤ 10% wCV (non-inferiority) | wCV = (SD of differences / mean) × 100; report 95% CI |  | Not started
3.2 | Calcified Plaque Volume | ≤ 20% wCV | wCV = (SD of differences / mean) × 100; report 95% CI |  | Not started
3.3 | Wall Volume | ≤ 30% wCV | wCV = (SD of differences / mean) × 100; report 95% CI |  | Not started
3.4 | Total Plaque Volume | ≤ 30% wCV | wCV = (SD of differences / mean) × 100; report 95% CI |  | Not started

## Gate 4 - Bias
Gate 4 — Systematic Bias & Agreement
Performance — Flag for clinical review on failure; may restrict use to specific indications
ID | Criterion | Threshold | Method | Owner | Status | Evidence Link / Location | Evidence Description | Date Completed | Notes
4.1 | Mean Bias — Lumen Volume | ¦bias¦ < 5% of mean lumen volume | Bland-Altman plot; report LoA |  | Not started
4.2 | Mean Bias — Calcified Plaque | ¦bias¦ < 10% of mean calc volume | Bland-Altman plot; proportional bias test |  | Not started
4.3 | Limits of Agreement | LoA within ±1.96 SD; no proportional bias (r² < 0.1) | Regress residuals on mean |  | Not started

## Advisory
Advisory — Contextual & Operational Checks
Inform labelling and use restrictions; do not block qualification
ID | Criterion | Threshold | Method | Owner | Status | Evidence Link / Location | Evidence Description | Date Completed | Notes
A.1 | HU Calibration Stability | Verify 130 HU threshold valid; map spectral to conventional HU if needed | Scanner calibration review |  | Not started
A.2 | Radiation Dose | Document DLP/CTDI; no dose increase that offsets clinical utility | Protocol dose comparison |  | Not started
A.3 | Patient Subgroup Stratification | Report wCV by calcium burden and BMI | Subgroup heterogeneity analysis |  | Not started
A.4 | Reader/Operator Variability | Inter-reader editing variability within expected range | Second analyst completes N = 10 Carolyn already did |  | Not started
A.5 | Compatibility with Ongoing Projects | No new failure modes with AVTE/AWAL | Regression testing |  | Not started
A.6 | Software Version Lock | Qualify against plaque algos ≥ Nov 2025; re-qualify on major updates | Version documentation |  | Not started