# PCCT Scanner Qualification — Statistical Method & Results

Gate 3 (quantitative reproducibility) — cross-scanner PCCT vs EID vs the B.1P Delta Validation OQ (730-CVV-040). Canonical vessel-overlap region, length-normalized, N = 25 paired patients. Two data vintages: v1 (original) and v2 (2026-07-07 re-work). Preliminary (target N ≥ 30).

# 1. Purpose & scope

Establish, on a like-for-like basis with the validated B.1P inter-operator reproducibility, whether cross-scanner (PCCT vs EID) quantitative reproducibility is non-inferior to the OQ limits. The systematic modality difference (bias) is separated from random dispersion and assessed in Gate 4; this document covers the Gate-3 reproducibility (wCV) method and the scanner-attributable variance decomposition.

# 2. Measurement model

For patient i, scanner s ∈ {PCCT, EID}, reader r, on the log(x+1), length-normalized scale:

Y_isr = μ + P_i + S_s + (PS)_is + R_r + e_isr

| Term | Meaning | Role |
| --- | --- | --- |
| P_i | patient (between-subject) effect | not part of within-subject reproducibility |
| S_s | scanner main effect (systematic modality bias) | FIXED → Gate 4 (not variance) |
| (PS)_is | patient × scanner interaction | σ²_scanner — the scanner-attributable variance |
| R_r | reader (inter-operator) effect | σ²_R — measured by the OQ |
| e_isr | residual / repeat error | σ²_e — measured by the OQ |

Only one read per patient·scanner is available, so only the paired difference d_i = Y_PCCT − Y_EID is observed. P_i cancels, giving d_i = Δ + [(PS) difference] + [R − R′] + [e − e′], hence Var(d)/2 = σ²_PS + σ²_R + σ²_e.

# 3. Estimators & acceptance

wCV (within-subject CV) — variance-component / random-effects estimator per Quan & Shih (1996), matching the OQ (730-CVV-040 §11: log(x+1) + linear mixed model). For paired (k=2) data the within-subject variance component is σ²_w = mean(d²)/2.

| Quantity | Formula | Purpose |
| --- | --- | --- |
| Systematic bias Δ | mean(d) (log scale) | Gate 4 (BA bias) |
| σ²_total | mean(d²)/2 = σ²_random + Δ²/2 | within-subject var incl. bias |
| σ²_random (scanner term) | Var(d)/2 | bias-removed dispersion (Gate 3 basis) |
| σ²_OQ,within | ln(1 + (OQ_logCV/100)²) | reader+repeat, imported from OQ |
| σ²_scanner | Var(d)/2 − σ²_OQ,within | scanner×patient interaction |
| wCV (log) | √(exp(σ²)−1)·100 | report/compare on log scale |
| wCV (untransf.) | √σ²_w / grand-mean · 100 | report on untransformed scale |

95% CIs: 2000-sample bootstrap over patients (seed 42). Acceptance — Gate 3: PCCT wCV 95% CI OVERLAPS the delta-OQ wCV CI (scanner-term basis). Scanner-attributable: (a) scanner wCV ≤ OQ inter-observer limit, and (b) σ²_scanner 95% CI includes 0 (not distinguishable from zero). The scanner-attributable subtraction is done on the log scale (there the CV↔variance relation is mean-independent, so the OQ term is transportable across studies).

# 4. OQ reference (730-CVV-040 / ATTACHMENT 2, patient level)

| Endpoint | OQ log-wCV [95% CI] | OQ untransf-wCV [95% CI] |
| --- | --- | --- |
| Lumen | 7.5% [5.9, 9.3] | 9.32% [7.09, 11.58] |
| Wall | 13.1% [10.2, 16.0] | 23.5% [17.88, 29.71] |
| Vessel | 8.7% [6.8, 10.6] | 10.18% [7.83, 12.38] |
| CALC | 4.5% [3.5, 5.4] | 25.78% [18.7, 35.75] |
| LRNC | 5.4% [4.1, 6.9] | 78.8% [48.09, 163.4] |
| NonCALC Matrix | 13.6% [10.8, 16.4] | 32.61% [24.87, 41.94] |
| Total Plaque | 13.2% [10.4, 16.0] | 27.08% [20.68, 34.74] |

# 5. Results — variance decomposition (log scale), SUB-SEGMENT (extent-matched)

v1 sub-seg  (N=25):

| Endpoint | Δ bias | σ²_total | σ²_rand | σ²_OQ | σ²_scan | wCV_tot | wCV_scanT | wCV_attrib |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Lumen | -0.089 | 0.01354 | 0.00996 | 0.00561 | 0.00435 | 11.7 | 10.0 | 6.6 |
| Wall | -0.058 | 0.01540 | 0.01431 | 0.01702 | -0.00270 | 12.5 | 12.0 | 0.0 |
| Vessel | -0.083 | 0.01126 | 0.00815 | 0.00754 | 0.00061 | 10.6 | 9.0 | 2.5 |
| CALC | +0.079 | 0.00577 | 0.00275 | 0.00202 | 0.00073 | 7.6 | 5.2 | 2.7 |
| LRNC | +0.043 | 0.00937 | 0.00879 | 0.00291 | 0.00588 | 9.7 | 9.4 | 7.7 |
| NonCALC Matrix | -0.167 | 0.04140 | 0.02866 | 0.01833 | 0.01033 | 20.6 | 17.1 | 10.2 |
| Total Plaque | -0.068 | 0.02274 | 0.02129 | 0.01727 | 0.00402 | 15.2 | 14.7 | 6.3 |

v2 sub-seg  (N=25):

| Endpoint | Δ bias | σ²_total | σ²_rand | σ²_OQ | σ²_scan | wCV_tot | wCV_scanT | wCV_attrib |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Lumen | -0.017 | 0.01014 | 0.01042 | 0.00561 | 0.00481 | 10.1 | 10.2 | 6.9 |
| Wall | -0.171 | 0.03077 | 0.01680 | 0.01702 | -0.00022 | 17.7 | 13.0 | 0.0 |
| Vessel | -0.072 | 0.01028 | 0.00803 | 0.00754 | 0.00049 | 10.2 | 9.0 | 2.2 |
| CALC | +0.002 | 0.00392 | 0.00408 | 0.00202 | 0.00206 | 6.3 | 6.4 | 4.5 |
| LRNC | +0.008 | 0.00380 | 0.00392 | 0.00291 | 0.00101 | 6.2 | 6.3 | 3.2 |
| NonCALC Matrix | -0.282 | 0.07169 | 0.03316 | 0.01833 | 0.01483 | 27.3 | 18.4 | 12.2 |
| Total Plaque | -0.224 | 0.04874 | 0.02467 | 0.01727 | 0.00740 | 22.3 | 15.8 | 8.6 |

Results are on the sub-segment (extent-matched) region for v1 and v2. σ²_random / σ²_scanner are small (the extent term is removed by the sub-segment intersection) and stable v1→v2 — the 07-07 re-work changed the systematic bias (Gate 4), not the random reproducibility (Gate 3). Gate 4 bias (§8) is assessed on the canonical length-normalized scale — the apples-to-apples match to the length-normalized OQ reference (the sub-segment is a separate extent-controlled view).

# 6. Gate 3 acceptance — scanner-term wCV vs OQ (95% CI overlap; P=pass/F=fail)

Log scale — process outputs (primary) + plaque:

| Endpoint | OQ wCV | v1 sub-seg | v2 sub-seg |
| --- | --- | --- | --- |
| Lumen | 7.5 [5.9,9.3] | 10.0 P | 10.2 P |
| Wall | 13.1 [10.2,16.0] | 12.0 P | 13.0 P |
| Vessel | 8.7 [6.8,10.6] | 9.0 P | 9.0 P |
| CALC | 4.5 [3.5,5.4] | 5.2 P | 6.4 P |
| LRNC | 5.4 [4.1,6.9] | 9.4 P | 6.3 P |
| NonCALC Matrix | 13.6 [10.8,16.4] | 17.1 P | 18.4 P |
| Total Plaque | 13.2 [10.4,16.0] | 14.7 P | 15.8 P |

Untransformed scale — plaque (reported on both scales):

| Endpoint | OQ wCV | v1 sub-seg | v2 sub-seg |
| --- | --- | --- | --- |
| CALC | 25.78 [18.7,35.75] | 18.6 P | 24.1 P |
| LRNC | 78.8 [48.09,163.4] | 109.6 P | 95.7 P |
| NonCALC Matrix | 32.61 [24.87,41.94] | 35.5 P | 45.7 P |
| Total Plaque | 27.08 [20.68,34.74] | 25.7 P | 35.8 P |

Cells: scanner-term wCV% and overlap verdict; 95% CIs are in §5 / gate3_comparison.csv.

# 7. Scanner-attributable variance vs OQ — SUB-SEGMENT (variance scale; a: σ²_scan ≤ σ²_OQ, b: CI∋0)

| Endpoint | σ²_OQ | v1 sub-seg σ²_scanner [95% CI] a/b | v2 sub-seg σ²_scanner [95% CI] a/b |
| --- | --- | --- | --- |
| Lumen | 0.0056 | +0.0043 [-0.0007,+0.0101] P/P | +0.0048 [-0.0012,+0.0113] P/P |
| Wall | 0.0170 | -0.0027 [-0.0091,+0.0028] P/P | -0.0002 [-0.0096,+0.0099] P/P |
| Vessel | 0.0075 | +0.0006 [-0.0027,+0.0037] P/P | +0.0005 [-0.0033,+0.0047] P/P |
| CALC | 0.0020 | +0.0007 [-0.0007,+0.0019] P/P | +0.0021 [-0.0008,+0.0051] F/P |
| LRNC | 0.0029 | +0.0059 [-0.0014,+0.0180] F/P | +0.0010 [-0.0021,+0.0060] P/P |
| NonCALC Matrix | 0.0183 | +0.0103 [-0.0024,+0.0240] P/P | +0.0148 [-0.0006,+0.0289] P/P |
| Total Plaque | 0.0173 | +0.0040 [-0.0052,+0.0128] P/P | +0.0074 [-0.0061,+0.0214] P/P |

Cells: scanner-attributable variance σ²_scanner = Var(d_log)/2 − σ²_OQ (log scale) with bootstrap 95% CI, on the extent-matched sub-segment. a: σ²_scanner ≤ σ²_OQ (P/F); b: CI includes 0 (P/F). On the variance scale the estimate and its CI lower bound can be NEGATIVE — a negative value / lower bound ≤ 0 means the residual cross-scanner variance is indistinguishable from (or below) the OQ reader+repeat variance. (The equivalent floored wCV% = √(exp(max(σ²,0))−1)×100 is in gate_results/scanner_attributable_subsegment.csv.) After removing reader+repeat (OQ) and traced extent (sub-segment), σ²_scanner is ≤ σ²_OQ and CI-includes-0 for nearly all endpoints.

# 8. Gate 4 — systematic bias (untransformed) vs OQ

The OQ Bland-Altman bias reference (730-CVV-040 ATTACHMENT 2, "Untransformed analysis") is a fixed, UNTRANSFORMED and LENGTH-NORMALIZED (per-mm) quantity, computed on the full traced vessel (reader-repeat, each read with its own traced extent). The apples-to-apples PCCT comparison is therefore the CANONICAL (full-vessel) length-normalized bias — the same processing as the OQ. (Length-normalization reflects each read's traced extent exactly as in the OQ; the sub-segment intersection is a different, extent-controlled basis the OQ did not use, so it appears only as a raw-mm³ reference.) OQ tabulates only the untransformed BA bias, so no transformed comparison is possible. Plaque endpoints.

8a. OQ comparison — canonical (full-vessel) length-normalized bias, per-mm. OQ 95% CI vs PCCT v1/v2; overlap P/F.

| Endpoint | OQ bias [95% CI] | v1 canonical bias [95% CI] | ov | v2 canonical bias [95% CI] | ov |
| --- | --- | --- | --- | --- | --- |
| CALC | +0.020 [-0.01, 0.05] | +0.017 [-0.075,+0.114] | P | -0.095 [-0.192,-0.004] | P |
| LRNC | -0.030 [-0.06, -0.01] | +0.037 [-0.021,+0.122] | P | -0.009 [-0.052,+0.043] | P |
| NonCALC Matrix | -0.070 [-0.19, 0.05] | -0.427 [-0.689,-0.189] | P | -0.626 [-0.876,-0.398] | F |
| Total Plaque | -0.050 [-0.19, 0.08] | -0.373 [-0.749,-0.018] | P | -0.730 [-1.106,-0.398] | F |

8b. Raw-mm³ bias reference (actual volume differences; not length-normalized). Canonical retains the extent differential; sub-segment is extent-matched.

| Endpoint | v1 canon (mm³) | v2 canon (mm³) | v1 sub-seg (mm³) | v2 sub-seg (mm³) |
| --- | --- | --- | --- | --- |
| CALC | +27.7 | -3.2 | +24.9 | +0.0 |
| LRNC | +5.6 | -4.2 | +6.5 | -1.4 |
| NonCALC Matrix | -65.2 | -108.9 | -76.6 | -111.2 |
| Total Plaque | -31.9 | -116.2 | -45.2 | -111.1 |

Note: on the OQ-comparable canonical length-normalized basis, CALC and LRNC overlap the OQ bias CI (v2), while NonCALC Matrix and Total Plaque do not — a real modality bias (PCCT lower). The bias grew v1→v2 (07-07 re-work). Length-normalized per-mm values carry each read's traced extent (as in the OQ); the raw-mm³ columns and the extent-matched sub-segment (§6/§7) disentangle composition from extent.

8c. Per-patient v1→v2 shift (untransformed & log). v1 = light, v2 = dark; grey arrows show each patient's shift; dashed = v1 mean bias, solid = v2 mean bias.

# 9. Interpretation

With the scanner term (systematic bias removed), the process outputs (Lumen, Wall, Vessel) pass the OQ CI-overlap on BOTH data vintages, canonical region — no sub-segment alignment required. The essential ingredient is separating the systematic scanner bias into Gate 4.

The variance decomposition shows σ²_random / σ²_scanner are vintage-stable; the 07-07 re-work increased the systematic bias Δ, not the random dispersion.

Plaque reproducibility depends on scale: on the untransformed scale (wide OQ CIs) all plaque endpoints overlap; on the log scale the tightest-limit component (CALC) is borderline and just fails on v2. Both scales are reported.

Scanner-attributable variance (Var(d)/2 − σ²_OQ) is ≤ the OQ limit / CI-includes-0 for most endpoints on the canonical region, and on the extent-controlled sub-segment region (regenerated on v2, N=25) it is ≤ OQ / CI-includes-0 for nearly all endpoints — i.e. after removing reader+repeat (OQ) and traced extent (sub-segment), the residual scanner-attributable variability is within the OQ envelope.

# 10. Caveats & future work

N = 25 canonical / 25 sub-segment, preliminary (target ≥ 30); passes are CI-overlap-driven with wide CIs.

Canonical region retains a traced-extent differential (PCCT traces ~+26% longer within shared vessels; Gate 2 REVIEW). The sub-segment (distance-from-ostium) intersection removes it; regenerated on v2 (2026-07-07) data, N=25 (PT-124 excluded — no PCCT∩EID vessel overlap; PT-136 recovered as two left targets, LAD+Diagonal and Circumflex+Marginal). Pipeline: scripts/subsegment/.

Scanner-attributable subtraction assumes the OQ reader+repeat variance is transportable to the PCCT setting, and that PCCT/EID reads are by independent readers (single read per patient·scanner; true scanner isolation would need replicate reads).

# References

730-CVV-040 v0.1 — B.1P Delta Validation OQ Report (+ ATTACHMENT 2, patient-level results).

4-B1P-033 v2.0 — B.1P Operational Qualification Report.

Quan H, Shih WJ (1996). Assessing reproducibility by the within-subject coefficient of variation with random effects models. Biometrics 52:1195-1203.

Shoukri MM, Elkum N, Walter SD (2006). Interval estimation and optimal design for the within-subject coefficient of variation. BMC Med Res Methodol 6:24.